## 05, Supplier quote.EN
ANNEX 5: SUPPLIER QUOTATION - LOT 1&2
HI'S REQUIREMENTS | [5 empty] | BIDDER'S RESPONSE
HI Billing Address: FÉDÉRATION HANDICAP INTERNATIONAL Room 207-208, E3 Trung Tu diplomatic compound, no.6 Dang Van Ngu, Dong Da, Hanoi, Vietnam. | [5 empty] | Supplier Name: ….................................................
Delivery Address: Bach Mai hospital, No.78 Giai Phong road, Phuong Mai, Dong Da district, Hanoi, Vietnam. | [5 empty] | Quotation Date: ….................................................
Contact Name for the delivery: Ms.HUE/HI Logistics
No. | Item | Description | Remark | QTY | Unit | Technical specifications | Brand | Model | Price | Amount | Life span
I. REHAB CENTER
1 | Air compression machine for full body | General requirement: - 100% new, manufactured from 2022 onwards - ISO 13485 quality standards or equivalent - Operating conditions: + Maximum operating temperature: ≥ 30 degrees Celsius + Maximum operating humidity: ≥ 70% Supply requirement: - Main device: 01 pce - Accessory set includes: + Leg cuff: 01 pair + Arm cuff: 01 pce + Waist cuff: 01 pce + Hose for leg cuff: 01 pce + Hose for arm cuff: 01 pce + Hose for waist cuff: 01 pce Technical specifications: - ≥10 modes program - Can skip pressure at any chamber; Pressure value of each chamber can be set independently; There is an emergency stop switch - Maximum pressure: 200mmHg + 20% - Adjustable time: 0- 99 minutes - Rate voltage: AC 220- 230V, 50~60Hz - Power consumption: ≥ 100 VA |  | 02 | set
2 | Whole body vibrator machine | Include 1 main machine 1 main machine weight: 55Kg Color control panel 2 handrails and pillars on both sides Full body vibrating table |  | 02 | piece
3 | Transcranial direct current stimulation (tDCS) | Include: 1 electrical stimulator, 1 set of wires, 2 electrodes, 2 electrode pads, 2 electrode fixing rubber bands, 1 electric charger: - Stimulator: 71mm x 94mm x 15mm; Weight: 66g - Program set: 71mm x 62mm x 15mm; Weight: 46g - Storage module: 71mm x 39mm x 15mm; Weight: 42g - Charging device: 71mm x 61mm x 15mm; Weight: 46g - One way direct electric current intensity: all from -2000 microAmpe to +2000 microAmpe - Electric current deviation/difference: 2% maximum - Hardware offset: +/-10 micro Ampe - Voltage limit: maximum +/-16V - Power consumption: 0.25W maximum - Operates continuously for 90 minutes - Graphical display, one pressing button | - The machine is indicated for stroke patients: related to upper limb movement, language disorders, insomnia, depression... combined with rehabilitation methods to support stroke patients. | 01 | piece
4 | Functional electrical stimulation (upper limbs) | Include: 1 multi-function electrode set with bluetooth connection, 1 screen to adjust the exercise/pratising program Power supply: 3.7V, 1400mAh battery Dimensions: 26mm x 76mm x 52mm Weight: 91grams Maximum load capacity: 5000 ohm Maximum output power: 1.5W Pulse type: symmetrical two-phase or 2 compensate capacitors Pulse intensity: resolution 0 – 60mA, 1mA (for load<5000) Pulse width: 150us – 300us, 10us resolution Pulse frequency: resolution 1- 40Hz 1Hz Maximum output voltage: 180V Maximum output frequency: 1KHz | Support the patient in practicing the functions of the wrist, hand, opening and closing the fingers | 01 | piece
5 | Functional electrical stimulation ((lower limbs) | Include: 1 multi-function electrode set with bluetooth connection, 1 screen to adjust the exercise/pratising program Power supply: 3.7V, 1400mAh battery Dimensions: 26mm x 76mm x 52mm Weight: 91grams Maximum load capacity: 5000 ohm Maximum output power: 1.5W Two modules Bluetooth 3.0 and bluetooth 4.0. dual mode Pulse type: symmetrical two-phase or 2 compensate capacitors Pulse intensity: resolution 0 – 60mA, 1mA (for load<5000) Pulse width: 150us – 300us, 10us resolution Pulse frequency: resolution 1- 40Hz 1Hz Maximum output voltage: 180V Maximum output frequency: 1KHz | Support patients with pratising gait /walking | 01 | piece
6 | Electrotherapy unit (treatment of patients with swallowing and voice disorders by stimulation currents) | 1 Main unit incl. at least standard accessories as below: Elastic fabric tape 1 pc; Hand button 1 pc; Power cable 1 pc; Patient cable 1 pc; Plate electrode 4 pcs; Viscose bag 4 pcs; Instructions for use (English-Vietnamese) 1 pc. -Microprocessor control -Number of treatment channels: 2 K12 channels. -Pronunciation training type: by low frequency and mid frequency. -Vocal stimulation electric currents: G, TENS, UR, IG30, IG50, FM, FaS, T/R, AMF, MT, KOTS -Output power: 20mA / 500 ohms ; G = 5mA / 500 ohms -Diagnostic program: check the patient's injury level -Machine control: - Multi-directional infinity knob; - Hand switch -Display of treatment settings: color touch type -Storage capacity: Yes, some treatment regimens can be saved by the user -Treatment combination tools: Yes, speakers play pronunciation and speaking exercises -Visibility: 3D images illustrate the area to be treated and electrode placement. -Automatically stops generating electric pulses when the output current at the electrode is interrupted. -Treatment time: from 0.5 to 60 minutes, adjustable level: 0.5 minutes. Support treatment results information on smartphones | The purpose of device is treatment of patients with swallowing and voice disorders by stimulation currents. | 01 | piece
7 | Biofeedback and neuromuscular electrical stimulation | Includes: 1 main machine, 1 set of electrode cables, 1 set of standard accessories Standard configuration: - Main machine with color touch screen: 01 machine - Electrode cables for dysphagia treatment: 04 pieces - Reference electrode cable: 01 piece - Switch: 01 piece - Shock-proof rubber case: 01 piece - Machine stand: 01 piece - Paste electrode: 01 package - Reference electrode: 01 - 1.5V AAA batteries: 04 pieces - User manual (Vietnamese): 01 book |  | 01 | piece
8 | TENS/EMS combo stimulation | Main machine: 01 pc Suitcase for carrying machine: 01 piece Electrode cable: 01 set Sticky electrode plate: 01 set Channel: dual, isolated between channels Pulse Amplitude: adjustable, 0-110mA peak into 500 ohm load each channel. Wave form: Asymmetrical Bi-Phasic square pulse. Voltage: 0-55V( load: 500ohm) Power source: One 9 Volt Battery Size: 11,8(L)x 6cm (W)x 3,1cm (H) Weight: 157grams with battery Pulse rate: adjustable from 2 to 150 Hz, 1Hz/ step Pulse Width: Adjustable from 50- 300 microseconds, 10 µs /step On Time: Adjustable 2-90 seconds, 1 Sec./step Off Time: Adjustable 2-90 seconds, 1 Sec./step Ramp Time: Adjustable 1-8 seconds, 1 Sec./ step, The "On" time will increase and decrease in the settting value. Mode: Six TENS modes: B(burst), N (normal), M (modulation rate & width), S1 (Modulation Width), S2 (Modulation Width) and P three EMS modes: C (constant), S (synchronous), A (Alternate) B( Burst mode): Burst rate: adjustable 0.5 ~ 5Hz, pulse width adjustable 50- 300 µs, Frequency fixed = 100Hz N( Normal mode): the pulse rate and pulse width are adjustable, it generates continous stimulation based on the setting value. M( modulation mode): modulation mode is a combination of pulse rate and pulse width modulation. S1 (S1 mode): pulse width is automaticaaly varied in a cyclic pattern over a nominal 10 second period. S2 (S2 mode): same as S1 but pulse width decrease over a period of 4 seconds from the initial setting to a value 70% less. C (Constant mode): The pulse rate and pulse width are adjustable. |  | 04 | piece
II. STROKE CENTER
1 | Operating Table | I. General requirements - 100% new equipment, manufactured from 2023 onwards. - Manufacturer meets ISO 13485 quality management standards - Power supply: 100-240V, 50/60Hz - Operating environment: + Maximum temperature: 35°C; + Maximum humidity: 70% II. Configuration requirements - Electro-hydraulic multi-function operating table and accessories: 01 set Which includes: + Body support set with cushion: 01 set; + Handheld control panel: 01 set; + Mattress: 01 set; + Anesthesia screen frame: 01 piece; + Arm board: 02 pieces; + Body fixation belt: 01 pc III. Technical Specifications - Electro-hydraulic multi-function operating table includes the following parts: head section, back section, pelvic section and leg section, leg section can be split Hand switch controller with functions + Electro-hydraulic control with control keys different positions + Reverse trendenburg / trendenburg + Table height up/down + Lateral tile left / right + Back section up/down. + On/Off (on the operating table leg) + Base cover, columns cover and 2 side rails are made of stainless steel, easy to clean and sterilize. + The mattress is thickness 5 cm and is made of waterproof material. + Caster: 4 casters (2 swivel caster, 2 fixed caster) + Mechanism locking function controlled by electro-hydraulic system + There is an emergency stop switch + Automatic return to level mechanism function + Table top size: 2063 x 500 mm + Adjustable height: 700 mm to 1000 mm + Trendelenburg/reverse trendenburg: 30 degrees + Lateral tilt (left / right): 20 degrees + Head section: Up 30 degrees, down 90 degrees, + Back section: Up 60º / Down 40º + Leg section bending: Up 0º, Down 90º, head and leg section can be swapped positions |  | 1 | ea
2 | Operation Light (Ceiling type) | I. General requirements - 100% new equipment, manufactured from 2023 onwards. - Manufacturer meets ISO 13485 quality management standards - Power supply: 100-240V, 50/60Hz - Operating environment: + Maximum temperature: 35°C + Maximum humidity: 70% II. Configuration requirements - Operating lamp with 2 reflectors: 01 set III. Technical Specifications - Small, compact lamp head and the lamp arm operates smoothly and gently - Using LED lights with high color rendering technology - Flexible control of light field - Uniform illumination at any position - Lighting by reflection method - Adjustment range light intensity 8 levels and control light field - Lamp structure: include 2 reflectors, with each reflector having: + Number of Led units: 4 + Number of Led devices: 72 + Lamp head diameter: 580mm + Maximum luminance: 140.000 lux + Irradiance (at 1m): 470W/m2 - Led service life: 40.000 hrs - Color temperature: 4100°K ± 150°K - Colore rendering index: Ra 94 - Intensity control: 8 steps ( 30 to 100%) on the wall panel or side panel - Light field and illuminance adjustment: center handle, side panel of lamp head and Operation wall panel |  | 1 | set
3 | NeuroSmart high speed motor system for crainial surgery | I. General requirements: - 100% new equipment, manufactured from 2023 onwards. - Manufacturer meets ISO 13485 quality management standards - Power supply: 110-240VAC; 50-60Hz - Operation conditions: + Temperature ≥ 27o C + Relative humidity ≥ 80% II. Configuration & technique requirements: - Control unit: 01 ea - Footswitch: 01 ea - High-speed motor: 01 cái - Perforator: 01 cái - Cranial perforator: 10 cái - Wash-rinse adapter: 01 cái - Instrument tray: 01 cái III. Technique requirements: 1. Central control machine body: - Power consumption ≥ 300 VA - Voltage: 100V- 240V; Frequency 50- 60 Hz - Integrated irrigation pump: 0- 100 ml/min - Weight ≤ 6kg - Maximum speed ≥ 75.000 rpm 2. Footswitch: - Waterproof standard: IPX8 - Length of cable ≥ 4m - Handle 3. High-speed motor: - High-speed motor with integrated cable ≥ 3m length, autoclavable, speed from 10.000-80.000 rpm, power supply ≥110 W 4. Perforator handpiece: - Speed about 1.000 rpm - Hudson chuck 5. Single-use cranial perforator: - Inner/outer diameter 11/14mm, for bone skull min 3mm thick 6. Wash-rinse adaptor: - 4 connection positon to wash 4 handpieces silmutaleously 7. Instrument tray: Instrument tray with lid and silicon holders, autoclavable, dimensions 482 x 246 x73 mm |  | 1 | system/set
4 | High frequency electrosurgical scalpel | I. General requirements ˗ Production year 2023 onwards. ˗ Voltage used: 220V/50 Hz ˗ Operating ambient temperature ranges from 10°C to 40°C ˗ Operating ambient humidity up to 80% II. Technical configuration ˗ Main machine: 01 machine ˗ Electrosurgical pencil: 01 Pcs ˗ Bipolar forceps Classic, straight: 01 Pcs ˗ Bipolar connecting cable: 01 Pcs ˗ Return electrode cable: 01 Pcs ˗ Neutral electrode glued: 50 Pcs ˗ One-pedal foot switch f: 01 Pcs ˗ Two-pedal foot switch f: 01 Pcs - Electric scalpel cart (bought in Vietnam): 01 Pcs ˗ User manual: 01 set III. Specifications: ˗ The machine is equipped with an automatic output power adjustment feature, providing enough voltage necessary for the selected cutting mode or controlling the electric arc current intensity, helping to protect tissue from unwanted heat damage. intentional ˗ Can be used in open surgery, endoscopy, lung surgery, dermatology,... ˗ There is a neutral socket that can be mounted to standard round and 2-pin standards ˗ Cutting power 200 Watt ˗ Coag power 120 Watt ˗ Power adjustment thanks to voltage regulation technology, for highly regenerative cuts ˗ The program sets up 9 user programs ˗ The machine is equipped with a Nessy safety system that monitors neutral electrode impedance ˗ There is an automatic mode to monitor the entire system when operating, displaying names and error codes on the screen and giving audible warnings. ˗ Cutting and hemostatic modes operate at current frequency: 350 kHz • The following cutting modes are available: ˗ AUTO CUT mode ˗ Maximum high-frequency voltage: from 740 Vp ˗ Maximum output power at rated impedance: 200W • The following Coag modes: ˗ Forced Coag + Maximum high-frequency voltage: 2000 Vp + Maximum output power at rated impedance: 120W ˗ Bipolar Coag: With 8 effects + Maximum high-frequency voltage: 200Vp + Maximum output power at rated impedance: 120W ˗ Soft Coag + Maximum high frequency voltage: 190Vp + Maximum output power at rated impedance: 120W |  | 1 | ea
5 | Patient Transport Stretcher 3 cranks | I. General requirements - 100% brand new, manufactured from 2023 onwards. - Manufacturer meets ISO 13485 quality management standards - Operating environment: + Maximum temperature: up to 40°C + Maximum humidity: up to 75% II. Configuration requirements - Frame with 04 locking casters: 01 set - Side rails on both sides: 01 set - Hand crank: 03 pcs - Mattress: 01 set - Infusion poles: 02 pcs - Oxygen tank holder: 01 pc III. Technical data - Multi-function patient stretcher trolley controlled by 3 cranks. - Allows control of the backest, knee, raising/lowering the stretcher height. - Material: stretcher is made of stainless steel 304 - The mattress base is made of stainless steel, divided into 4 parts: a head rest, a hip section, a knee section and a lower leg rest - There are side rails on both sides, able to be folded down when needed and have lock switch. - Equipped with storage tray placed below the stretcher - Equipped with an oxygen tank holder - Surround with the rubber bumper to help prevent impact during movement. - Mattress is made of PU foam material, mattress cover is artificial leather, 50 mm thickness - Overall length: 1900 mm - Overall width: 750mm - Backrest adjustment: 0 – 80° - Knee rest adjustment: 0 – 40° - Height adjustment: 520 - 750 mm - The lower leg rest can be raised and lowered by hand - The bed is equipped with 3 lockable casters and 1 steeing lock caster - Casters diameter: 150mm - Maximum loading capacity: 200 kg |  | 2 | set
| Price | Door to Door including insurance, installation and training.
| Currency | USD
| Delivery lead time | ( to be specified by the supplier)
| Warranty | ( to be specified by the supplier)
| Other services after-sales | ( to be specified by the supplier)
| Payment terms | ( to be specified by the supplier)
| Payment modality | Bank transfer or check
| Tender validility | 60 days
Other…..
Other…..
Other…..
* The bidders can also attach your bid price using the company’s own template (Proposal/Quotation) and letterhead, however, all the required information must be addressed. * If any discrepancy information between the Supplier Quotation and your proposal/quotation, the Supplier Quotation will be valid
SUPPLIER Name, Signature and Stamp

## 05, Supplier quote.VN
ANNEX 5: SUPPLIER QUOTATION - LOT 1&2
HI'S REQUIREMENTS | [5 empty] | BIDDER'S RESPONSE
HI Billing Address: FÉDÉRATION HANDICAP INTERNATIONAL Room 207-208, E3 Trung Tu diplomatic compound, no.6 Dang Van Ngu, Dong Da, Hanoi, Vietnam. | [5 empty] | Supplier Name: ….................................................
Delivery Address: Bach Mai hospital, No.78 Giai Phong road, Phuong Mai, Dong Da district, Hanoi, Vietnam. | [5 empty] | Quotation Date: ….................................................
Contact Name for the delivery: Ms.HUE/HI Logistics
No. | Item | Description | Remark | QTY | Unit | Technical specifications | Brand | Model | Price | Amount | Life span
I. REHAB CENTER
1 | Máy nén bóp áp lực hơi toàn thân | A.Yêu cầu chung - Thiết bị mới 100%, sản xuất từ 2022 trở về sau - Thiết bị đạt tiêu chuẩn chất lượng ISO 13485 hoặc tương đương - Điều kiện hoạt động: + Nhiệt độ hoạt động tối đa tới: ≥ 30 độ C + Độ ẩm hoạt động tối đa tới: ≥ 70% B. Yêu cầu cấu hình Máy chính : 01 cái Bộ phụ kiện bao gồm: Ống nén khí xoa bóp chân : 01 đôi Ống nén khí xoa bóp tay : 01 cái Ống nén khí xoa bóp vùng lưng : 01 cái Dây dẫn khí cho chân : 01 cái Dây dẫn khí cho tay : 01 cái Dây dẫn khí cho lưng : 01 cái C. Chỉ tiêu kỹ thuật ≥10 chương trình chế độ Có thể bỏ nén áp lực tại khoang khí bất kỳ Có thể cài đặt nén áp lực từng khoang Có công tắc dừng khẩn cấp Áp suất tối đa: 200mmHg + 20% Thời gian điều trị: 0- 99 phút, điều chỉnh được Nguồn điện: AC 220- 230V, 50~60Hz Công suất tiêu thụ: ≥ 100 VA |  | 02 | bộ
2 | Máy rung toàn thân | 1 máy chính trọng lượng: 55Kg Bảng điều khiển màu 2 tay vịn và trụ giữ 2 bên Bàn rung toàn thân |  | 02 | chiếc
3 | Máy kích thích điện 1 chiều xuyên sọ | Bao gồm: 1 máy kích thích điện, 1 bộ dây dẫn, 2 điện cực, 2 miếng đệm điện cực, 2 dây cao su cố định điẹn cực, 1 sạc điện Bộ kích thích: 71mm x 94mm x 15mm; Trọng lượng: 66g Bộ chương trình: 71mm x 62mm x 15mm; Trọng lượng: 46g Mô-đun lưu trữ: 71mm x 39mm x 15mm; Trọng lượng: 42g Thiết bị sạc: 71mm x 61mm x 15mm; Trọng lượng: 46g Cường độ dòng điện một chiều đều: từ -2000microAmpe đến +2000microAmpe Độ lệch dòng điện: tối đa 2% Offset phần cứng: +/-10 microAmpe Giới hạn điện thế: tối đa +/-16V Tiêu thụ năng lượng: tối đa 0.25W Hoạt động liên tục 90 phút Hiển thị đồ hoạ, một nút ấn | Máy được chỉ định cho đối tượng người bệnh đột quỵ não: liên qiuan đến vận động chi trên, rối loạn ngôn ngữ, mất ngủ , trầm cảm… kết hợp với các phương pháp phục hồi chức năng hỗ trợ cho người bệnh đột quỵ não . | 01 | chiếc
4 | Máy kích thích điện chức năng (chi trên) | Bao gồm: 1 bộ điện cực đa năng có kết nối bluetooth, 1 màn hình điều chỉnh chương trình tập luyện Nguồn điện: Pin 3,7V, 1400mAh Kích thước: 26mm x 76mm x 52mm Trọng lượng: 91gram Công suất tải tối đa: 5000 ohm Công suất đầu ra tối đa: 1,5W Loại xung: hai pha đối xứng hoặc hai pha bù Cường độ xung: độ phân giải 0 – 60mA, 1mA (đối với tải<5000) Độ rộng xung: 150us – 300us, độ phân giải 10us Tần số xung: độ phân giải 1- 40Hz 1Hz Điện áp đầu ra tối đa: 180V Tần số đầu ra tối đa: 1KHz | Hỗ trợ người bệnh tập luyện các chức năng của cổ tay, bàn tay, mở và nắm các ngón tay | 01 | chiếc
5 | Máy kích thích điện chức năng (chi dưới) | Bao gồm: 1 bộ điện cực đa năng có kết nối bluetooth, 1 màn hình điều chỉnh chương trình tập luyện Nguồn điện: Pin 3,7V, 1400mAh Kích thước: 26mm x 76mm x 52mm Trọng lượng: 91gram Công suất tải tối đa: 5000 ohm Công suất đầu ra tối đa: 1,5W Hai mô-đun Bluetooth 3.0 và bluetooth 4.0. chế độ kép Loại xung: hai pha đối xứng hoặc hai pha bù Cường độ xung: độ phân giải 0 – 60mA, 1mA (đối với tải<5000) Độ rộng xung: 150us – 300us, độ phân giải 10us Tần số xung: độ phân giải 1- 40Hz 1Hz Điện áp đầu ra tối đa: 180V Tần số đầu ra tối đa: 1KHz | Hỗ trợ người bệnh tập luyện dáng đi | 01 | chiếc
6 | Máy điện thị liệu (Máy điều trị và kích thích phát âm) | Máy chính : 01 cái, kèm phụ kiện chuẩn sau: Đai thắt điện cực điều trị 01 cái; Công tắc tay 01 cái; Cáp nguồn cung cấp 01 Chiếc; Cáp nối điện cực 01 cái; Điện cực điều trị 04 cái; Đệm điện cực 04 cái; Tài liệu HDSD bảo dưỡng tiếng Anh+ tiếng Việt 01 bộ Điều khiển vi xử lý Số kênh điều trị: 2 kênh.K12 Kiểu luyện tập phát âm: bằng dòng điện tần số thấp và tần số trung. Các dòng điện kích thích phát âm: G, TENS, UR, IG30, IG50, FM, FaS, T/R, AMF, MT, KOTS Công suất phát: 20mA / 500 ohms ; G = 5mA / 500 ohms Chương trình chẩn đoán: kiểm tra mức độ tổn thương của bệnh nhân Điều khiển máy: - Núm xoay vô cực đa hướng - Công tắc tay Hiển thị các thông số cài đặt khi điều trị: kiểu cảm ứng màu Khả năng lưu trữ: Có, một số phác đồ điều trị do người sử dụng lưu vào Dụng cụ kết hợp điều trị: Có, loa phát tiếng các bài luyện âm, luyện nói Khả năng hiển thị:Hình ảnh 3D minh hoạ vùng sẽ điều trị và vị trí đặt điện cực. Tự động dừng phát xung điện khi dòng ra ở điện cực bị gián đoạn. Thời gian điều trị: từ 0,5 đến 60 phút, nấc chỉnh 0,5 phút. Hỗ trợ thông tin kết quả điều trị trên điện thoại thông minh | Hỗ trợ chẩn đoán và điều trị người bệnh phát âm nói và tập nuốt | 01 | chiếc
7 | Máy Biofeedback tập nuốt cầm tay | Bao gồm: 1 máy chính, 1 bộ cáp điện cực, 1 bộ phụ kiện chuẩn Cấu hình tiêu chuẩn: - Máy chính màn hình màu cảm ứng: 01 máy - Cáp điện cực điều trị khó nuốt: 04 cái - Cáp điện cực tham chiếu: 01 cái - Công tắc: 01 cái - Vỏ cao su chống sốc: 01 cái - Đế kê máy: 01 cái - Điện cực dán: 01 gói - Điện cực tham chiếu: 01 - Pin AAA 1,5V: 04 viên - Hướng dẫn sử dụng Tiếng Việt: 01 quyển |  | 01 | chiếc
8 | Máy kích thích TENS/ EMS cầm tay | Máy chính: 01 cái Vali đựng máy: 01 cái Cáp điện cực: 01 bộ Tấm điện cực dán: 01 bộ Số kênh: 2 kênh, điều chỉnh độc lập giữa hai kênh Xung: có thể điều chỉnh từ 0-110mA đỉnh 500 ohm mỗi kênh Dạng sóng: Xung vuông Bi-Phasic không đối xứng. Hiệu điện thế: 0-55V( tải 500ohm) Nguồn cấp: Pin 9V Kích thước: 11,8x 6x 3,1cm Trọng lượng: 157g bao gồm cả pin Tỉ lệ xung: Điều chỉnh 2-150 Hz, 1Hz mỗi bước Độ rộng xung: Điều chỉnh từ 50- 300 micro giây, chỉnh 10 micro giây mỗi bước Thời gian hoạt động: Điều chỉnh 2-90 giây, 1 giây mỗi bước chỉnh Ramp Time: chỉnh 1-8 giây, 1 giây mỗi bước, tăng và giảm trong cài đặt giá trị Kiểu dòng điện: 6 kiểu dòng TEN: B, N, M, S1, S2 và P 3 kiểu dòng EMS: C, S, A B( Burst mode): Tỉ số Burst điều chỉnh 0.5 đến 5Hz, độ rộng xung 50- 300 µs, tần số cố định 100Hz N( Normal mode): tỉ số và độ rộng xung điều chỉnh được, chạy liên tục dựa trên thông số cài đặt M( modulation mode): Kiểu kết hợp điều biến tỉ lệ và độ rộng xung. |  | 04 | chiếc
II. STROKE CENTER
1 | Bàn mổ đa năng | I. Yêu cầu chung - Thiết bị mới 100%, được sản xuất từ năm 2023 trở về sau. - Nhà sản xuất đạt tiêu chuẩn quản lý chất lượng ISO 13485 - Nguồn điện sử dụng: 100-220V, 50/60Hz - Môi trường hoạt động: + Nhiệt độ tối đa đến: 30°C + Độ ẩm tối đa đến: 75% II. Yêu cầu về cấu hình - Bàn mổ đa năng điện thủy lực và phụ kiện kèm theo: 01 bộ Trong đó đã bao gồm: + Bộ đỡ thân kèm đệm: 01 bộ + Bảng điều khiển cầm tay: 01 bộ + Đệm bàn: 01 bộ + Khung chắn gây mê: 01 chiếc + Đỡ tay có đệm: 02 cái + Đỡ chân có đệm: 02 cái + Đai cố định thân: 01 chiếc III. Chỉ tiêu kỹ thuật - Bàn mổ đa năng điện thủy lực gồm các bộ phận: Đỡ đầu, lưng, hông và chân 2 tấm tách rời nhau được Có bộ điều khiển cầm tay kéo dài với các chức năng: + Điều chỉnh điện thủy lực bàn bằng phím điều khiển cho các tư thế: + Điều khiển dốc bàn ngược / xuôi + Điều khiển nâng / hạ mặt bàn + Nghiêng bàn hai bên trái / phải + Nâng / hạ phần lưng. + Bật / Tắt (trên chân bàn mổ) + Chân, đế, cột và 2 thanh ray bên cạnh mép bàn được làm bằng thép không gỉ, dễ vệ sinh và tiệt trùng. + Đệm của bàn dày 5 cm được làm bằng vật liệu không thấm nước. + Bánh xe: 4 bánh xe (trong đó 2 bánh xoay, 2 bánh cố định) + Chức năng khóa cơ học bàn mổ + Có nút dừng khẩn cấp + Có chức năng tự động điều chỉnh về vị trí nằm ngang + Kích thước mặt bàn: 2063 x 500 mm + Độ cao điều chỉnh được: 700 mm đến 1000 mm + Trendelenburg (lên / xuống): 30 độ (Góc nghiêng theo phương dọc bàn mổ) + Độ nghiêng 2 bên bàn mổ (trái / phải): 20 độ + Góc nâng của tấm đầu: Lên 30 độ, xuống 90 độ, + Góc nâng/ hạ của tấm lưng: Lên 60º / Xuống 40º + Góc nâng của tấm chân: Lên 0º, Xuống 90º, tấm đỡ đầu và đỡ chân có thể tráo đổi vị trí cho nhau |  | 1 | chiếc
2 | Đèn mổ treo trần (2 chóa) | I. Yêu cầu chung - Thiết bị mới 100%, được sản xuất từ năm 2023 trở về sau. - Nhà sản xuất đạt tiêu chuẩn quản lý chất lượng ISO 13485 - Nguồn điện sử dụng: 100-240V, 50/60Hz - Môi trường hoạt động: + Nhiệt độ tối đa đến: 35°C + Độ ẩm tối đa đến: 70% II. Cấu hình kỹ thuật: - Đèn mổ 2 chóa: 01 bộ III. Chỉ tiêu kỹ thuật - Đầu đèn nhỏ, gọn và cánh tay đèn hoạt động trơn tru, nhẹ nhàng - Sử dụng đèn LED công nghệ hoàn màu siêu cao - Kiểm soát trường sáng linh hoạt - Chiếu sáng đồng nhất ở mọi vị trí - Chiếu sáng bằng phương pháp phản chiếu - Có thể điều chỉnh 8 mức chiếu sáng và kiểm soát trường sáng - Cấu trúc đèn: gồm 2 chóa, với mỗi chóa có: + Số khối LED: 4 + Số lượng đèn LED: 72 + Đường kính đầu đèn: 580mm + Cường độ sáng tối đa: 140.000lux + Năng lượng chiếu sáng (tại vị trí 1m): 470W/m2. - Tuổi thọ đèn Led: 40.000h - Nhiệt độ màu: 4,100K ± 150K - Chỉ số màu tạo ra: Ra 94 - Điều chỉnh cường độ sáng: 8 bước ( 30-100%) trên bảng điều khiển vách tường, và bên cạnh đầu đèn - Điều chỉnh trường sáng và cường độ sáng: tay cầm trung tâm, bảng điều khiển trên đầu đèn và bảng điều khiển gắn tường |  | 1 | bộ
3 | Hệ thống máy khoan tốc độ cao dùng trong phẫu thuật sọ não | I. Yêu cầu chung - Thiết bị mới 100%, được sản xuất từ năm 2023 trở về sau. - Nhà sản xuất đạt tiêu chuẩn quản lý chất lượng ISO 13485. - Nguồn điện sử dụng: 100-240VAC, 50/60Hz - Môi trường hoạt động: + Nhiệt độ tối đa đến: ≥ 27o C + Độ ẩm tối đa đến: ≥ 80% II. Yêu cầu cấu hình và kỹ thuật: - Thân máy điều khiển trung tâm: 01 cái - Bàn đạp điều khiển tốc độ: 01 cái - Mô tơ tốc độ cao: 01 cái - Tay khoan sọ não: 01 cái - Mũi khoan sọ não tự dừng: 10 cái - Dụng cụ làm sạch: 01 cái - Khay đựng và bảo quản: 01 cái III. Yêu cầu kỹ thuật: 1. Thân máy điều khiển trung tâm: - Công suất: ≥ 300 VA - Điện nguồn: 100V đến 240V; tần số từ 50 đến 60 Hz - Tích hợp bơm tưới rửa từ 0 đến 100 ml/ phút - Khối lượng: ≤ 6kg - Tốc độ tối đa ≥ 75.000 vòng/phút 2. Bàn đạp điều khiển tốc độ: - Chống thấm IPX8 - Dây cáp dài ≥ 4m - Có tay cầm 3. Mô tơ tốc độ cao: - Mô tơ tốc độ cao kèm dây cáp dài ≥ 3m, có thể hấp tiệt khuẩn, tốc độ từ 10.000-80.000 vòng/ phút, công suất ≥110 W 4. Tay khoan sọ não tự dừng: -Tốc độ khoan đạt đến 1.000 vòng/phút. - Chuôi lắp mũi khoan kiểu Hudson. 5. Mũi khoan sọ não tự dừng, dùng một lần: - Đường kính trong/ngoài: 11/14mm, dùng cho xương sọ độ dày 3mm 6. Dụng cụ làm sạch: - Có 4 cổng để có thể rửa đồng thời 4 tay/đầu chụp một lúc 7. Khay đựng vào bảo quản máy khoan: - Khay đựng và bảo quản máy khoan với nắp và các gá silicon, có thể hấp ẩm được, kích thước khoảng 482 x 246 x73 mm |  | 1 | hệ thống/ bộ
4 | Dao mổ điện cao tần | I. Yêu cầu chung ˗ Năm sản xuất 2023 trở về sau ˗ Điện áp sử dụng: 220V/50 Hz ˗ Nhiệt độ môi trường hoạt động: từ 10°C đến 40°C ˗ Độ ẩm môi trường hoạt động đến 80% II. Cấu hình kỹ thuật ˗ Máy chính: 01 máy ˗ Tay dao cắt đốt, sử dụng nhiều lần: 01 cái ˗ Kẹp lưỡng cực: 01 cái ˗ Dây nối lưỡng cực: 01 cái ˗ Dây nối trung tính: 01 cái ˗ Điện cực trung tính dán: 50 cái ˗ Bàn đạp điều khiển 1 nút bấm: 01 cái ˗ Bàn đạp điều khiển 2 nút bấm: 01 cái ˗ Xe đặt dao mổ điện (mua tại Việt Nam): 01 cái ˗ Tài liệu hướng dẫn sử dụng tiếng Anh, tiếng Việt: 01 bộ III. Thông số kỹ thuật: ˗ Máy được trang bị tính năng tự động điều chỉnh công suất đầu ra, cung cấp đủ điện áp cần thiết cho chế độ cắt đã lựa chọn hoặc kiểm soát cường độ dòng hồ quang điện, giúp bảo vệ mô tránh các tổn thương nhiệt không chủ định ˗ Có khả năng dùng trong các phẫu thuật mổ mở, nội soi, phổi, da liễu,... ˗ Có ổ cắm trung tính gắn được chuẩn tròn và chuẩn 2-pin ˗ Công suất cắt 200 Watt ˗ Công suất cầm máu 120 Watt ˗ Điều chỉnh công suất nhờ vào công nghệ điều chỉnh điện áp, cho vết cắt có khả năng tái sinh cao ˗ Chương trình lập cho người sử dụng 9 chương trình ˗ Máy được trang bị hệ thống an toàn theo dõi trở kháng điện cực trung tính ˗ Có chế độ tự động giám sát toàn bộ hệ thống khi hoạt động, hiển thị tên, mã lỗi trên màn hình đồng thời cảnh báo bằng âm thanh ˗ Các chế độ cắt và cầm máu hoạt động ở tần số dòng điện: 350 kHz • Có các chế độ cắt sau: ˗ Chế độ AUTO CUT ˗ Hiệu điện thế cao tần lớn nhất: từ 740 Vp ˗ Công suất đầu ra cực đại ở trở kháng định mức: 200W • Các chế độ cầm máu sau: ˗ Chế độ cầm máu cưỡng bức + Hiệu điện thế cao tần lớn nhất: 2000 Vp + Công suất đầu ra cực đại ở trở kháng định mức: 120W ˗ Chế độ cầm máu công suất thấp lưỡng cực: Với 8 hiệu ứng + Hiệu điện thế cao tần lớn nhất: 200Vp + Công suất đầu ra cực đại ở trở kháng định mức: 120W ˗ Chế độ cầm máu mềm + Hiệu điện thế cao tần lớn nhất: 190Vp + Công suất đầu ra cực đại ở trở kháng định mức: 120W |  | 1 | chiếc
5 | Cáng vận chuyển bệnh nhân 3 tay quay | I. Yêu cầu chung - Thiết bị mới 100%, được sản xuất từ năm 2023 trở về sau. - Nhà sản xuất đạt tiêu chuẩn quản lý chất lượng ISO 13485 - Môi trường hoạt động: + Nhiệt độ tối đa đến: 40°C + Độ ẩm tối đa đến: 75% II. Yêu cầu về cấu hình - Khung xe với 04 bánh xe có khóa: 01 bộ - Bộ tay vịn hai bên: 01 bộ - Tay quay: 03 cái - Đệm: 01 bộ - Cọc truyền: 02 cái - Giá đựng bình oxy: 01 chiếc III. Chỉ tiêu kỹ thuật - Xe đẩy cáng bệnh nhân đa năng điều khiển 3 tay quay. - Cho phép điều khiển phần lưng, khớp gối, nâng/ hạ chiều cao cáng. - Vật liệu: cáng được làm bằng vật liệu thép không gỉ 304 - Mặt nằm được làm bằng thép không gỉ, chia làm 4 phần: phần đầu, phần hông, phần gối và phần chân - Có tay vịn hai bên, cho phép gấp gọn xuống khi cần thiết và có khóa chốt - Có khay chứa đồ đặt bên dưới cáng - Có giá đỡ bình oxy - Có đệm bằng cao su phủ quanh cáng giúp chống va đập trong quá trình di chuyển. - Đệm được làm bằng vật liệu PU foam, vỏ đệm được bọc da nhân tạo, độ dày: 50 mm - Chiều dài cáng: 1900 mm - Chiều rộng cáng: 750mm - Góc nâng phần lưng: 0 – 80° - Góc nâng khớp gối: 0 – 40° - Điều chỉnh chiều cao trong khoảng: 520 – 750 mm - Cho phép nâng hạ phần chân dưới bằng tay - Ba bánh xe có khóa hãm, bánh xe còn lại giúp dẫn hướng - Đường kính bánh xe: 150mm - Tải trọng tối đa: 200 kg IV. YÊU CẦU KHÁC - Thời gian giao hàng: ≤ 120 ngày - Địa điểm giao hàng: Bệnh viện Bạch Mai Hà Nội 78 Giải Phóng, Phương Mai, Đống Đa, Hà Nội - Thời gian bảo hành: ≥ 24 tháng kể từ khi ký kết biên bản bàn giao, nghiệm thu với bên mua, Bên bán phải thực hiện bảo trì/hiệu chuẩn, theo quy định của nhà sản xuất. |  | 2 | bộ
| Price | Door to Door including insurance, installation and training.
| Currency | USD
| Delivery lead time | ( to be specified by the supplier)
| Warranty | ( to be specified by the supplier)
| Other services after-sales | ( to be specified by the supplier)
| Payment terms | ( to be specified by the supplier)
| Payment modality | Bank transfer or check
| Tender validility | 60 days
Other…..
Other…..
Other…..
* The bidders can also attach your bid price using the company’s own template (Proposal/Quotation) and letterhead, however, all the required information must be addressed. * If any discrepancy information between the Supplier Quotation and your proposal/quotation, the Supplier Quotation will be valid
SUPPLIER Name, Signature and Stamp